KING COUNTY ENVIRONMENTAL LABORATORY

QUALITY ASSURANCE REVIEW

for

DUWAMISH ESTUARY INVERTEBRATE TISSUE STUDY

RWSP WATER QUALITY ASSESSMENT

October 20, 1997

King County Environmental Laboratory

322 West Ewing Street

Seattle, Washington 98119-1507

KING COUNTY ENVIRONMENTAL LABORATORY

QUALITY ASSURANCE REVIEW

for

DUWAMISH ESTUARY INVERTEBRATE TISSUE STUDY

RWSP WATER QUALITY ASSESSMENT

Prepared by:

__________________________________

Scott J. Mickelson

Laboratory Project Manager

October 20, 1997

King County Environmental Laboratory

322 West Ewing Street

Seattle, Washington 98119-1507

INTRODUCTION

This Quality Assurance (QA) review accompanies analytical data submitted in connection with an invertebrate tissue study undertaken as part of the RWSP Water Quality Assessment.  The QA review is organized into the four sections listed below.

- General Comments

- Conventionals Chemistry

- Metals Chemistry

- Organics Chemistry

An overview of the approach used for the QA review is detailed in the General Comments section.  Additional information specific to each analysis is included in the appropriate analytical section. Quality control (QC) data are included in appendices for conventionals, metals and organics analyses.

GENERAL COMMENTS

Scope of Samples Submitted

This QA review is associated with invertebrate tissue samples collected in June, 1997.  These samples were comprised of invertebrates (mainly amphipods and other small arthropods) collected from the Duwamish River near Kellogg Island and a reference station in the Nisqually Delta.  The samples collected and the proposed analytical scheme are summarized in Table 1.  All analyses were performed by the King County Environmental Laboratory (KCEL).

Completeness

Completeness has been evaluated for this data submission and QA review by considering the following criteria:

- Comparing available data with the planned project analytical scheme summarized in Table 1.

- Compliance with storage conditions and holding times.

- Compliance with the complete set of quality control (QC) samples outlined in Table 2.

Methods

Analytical methods are noted in the applicable analytical sections of this QA review.

Target Lists

The reported target lists are noted in the applicable analytical sections of this QA review.

Detection Limits

The KCEL distinguishes between the Reporting Detection Limit (RDL) and the Method Detection Limit (MDL).

- The RDL is defined as the minimum concentration of a chemical constituent that can be reliably quantified.

- The MDL is defined as the minimum concentration of a chemical constituent that can be detected.

Storage Conditions and Holding Times

Storage conditions and holding times have been evaluated using guidelines established during the recent update of the Puget Sound Protocols (Recommended Guidelines for Sampling Marine Sediment, Water Column, and Tissue in Puget Sound - PSEP, 1996).

Method Blanks

Method blank results have been evaluated for the presence of positive analyte results at or greater than the MDL indicating possible laboratory contamination of samples.

Spike Blanks

Spike blank results have been evaluated to assess the effect of low or high recoveries on associated sample data.  Spike blank recoveries should not be used alone as an indicator of data quality for individual compounds.  These data should be reviewed in conjunction with matrix spike recoveries and standard reference material recoveries to evaluate the sum of method performance and possible matrix interferences for each compound.  Those compounds for which two or more of the recoveries are either above or below recommended QC limits should be reviewed to determine if associated sample data may be low or high biased.  A low recovery indicates that associated sample data may be low biased for that compound or element. A high recovery indicates that associated sample data may be high biased for that compound or element.

Surrogate Recoveries (Organics Only)

Surrogate recoveries have been evaluated to assess analytical bias on a sample by sample basis for a group of compounds such as BNAs or PCBs. A low surrogate recovery may indicate low biased sample data for a particular group of compounds and a high surrogate recovery may indicate high biased sample data, based on method performance and matrix interference.

Standard Reference Material

Standard reference material (SRM) results have been evaluated to assess the effect of low or high recoveries on associated sample data for a particular compound or element.  SRM recoveries should be reviewed in conjunction with spike blank and matrix spike recoveries.  A low SRM recovery may indicate low biased sample data for a particular element or compound and a high SRM recovery may indicate high biased sample data, based on method performance.

Matrix Spikes and Matrix Spike Duplicates

Matrix spike and matrix spike duplicate (MS/MSD) results have been evaluated to assess the effect of low or high recoveries on associated sample data.  These data should be evaluated in conjunction with spike blank and SRM data to assess low or high biased data for individual compounds or elements.  MS/MSD results will also be evaluated against each other to assess the precision of the analytical method in addition to matrix interferences.

Laboratory Duplicate Samples

Laboratory duplicate samples have been evaluated to assess the precision of the analytical method.  However, not all duplicate results have been used as an indicator of data quality.  Only sets of duplicate results which contain at least one result significantly greater than the MDL have been considered for data qualification.  Where an RDL is present, only duplicate data that contain at least one result greater than the RDL have been considered for data qualification.  These guidelines have been used to account for the fact that precision obtained near the detection limit is not representative of precision obtained throughout the entire analytical range.

Data Qualifiers

The current policy of the KCEL is to only qualify data from tissue analyses based upon method blank contamination and non-compliance with sample storage or holding time criteria.  Data qualified for blank contamination is noted with a B flag.  Data qualified for storage condition or holding time issues is noted with an H flag.  It is up to the data user to qualify data based upon other QC criteria.

Units and Significant Figures

Data have been reported in accordance with laboratory policy at the time of data generation.  When an RDL and MDL are reported, data have been reported to three significant figures above the RDL, and two significant figures equal to or below the RDL.

METALS

Completeness

Metals data are reported for samples 11193-1, 11193-2, 11444-1, and 11444-2.  These samples were analyzed for mercury and other metals in association with the complete set of QC samples outlined in Table 2.

Methods

Mercury analysis was performed in accordance with EPA Method 7471 (CVAA).  Chromium and zinc analyses were performed in accordance with EPA Method 6010 (ICP).  All other metals analyses were performed in accordance with EPA Method 6020 (ICPMS).  Sample digestion was performed following guidelines suggested in Recommended Guidelines for Measuring Metals in Puget Sound Marine Water, Sediment and Tissue Samples (PSEP, 1996).

Target List

The reported target list includes antimony, arsenic, cadmium, chromium, copper, lead, mercury, nickel, silver, and zinc.

Detection Limits, Units, and Significant Figures

Data are reported in accordance with laboratory policy at the time the data were generated.  A positive result and/or MDL and RDL have been reported for all metals.  Sample results are reported in units of mg/Kg on a wet weight basis.  Data are reported to three significant figures for results greater than the RDL and two significant figures for results equal to or less than the RDL.  For results reported with less than two or three significant figures, significant zeroes are implied.

Storage Conditions and Holding Times

Sample storage conditions and holding times have been evaluated using guidelines established during the recent update of the Puget Sound Protocols (Recommended Guidelines for Sampling Marine Sediment, Water Column, and Tissue in Puget Sound - PSEP, 1996).  The criteria used to evaluate storage conditions and holding times for metals analyses are listed in the table below.

| Parameter | Frozen Holding Time | Refrigerated Holding Time |
| --- | --- | --- |
| Mercury | 28 Days | Not Recommended |
| Metals | 2 Years | Not Recommended |

All tissue samples in this data submission were stored frozen, both prior and subsequent to processing and homogenization.  All analyses were completed within holding times.

Method Blanks

All mercury and metals method blank results were less than the MDL.

Spike Blanks

All mercury and metals spike blank recoveries were within the recommended QC limits of 80 to 120%.

Standard Reference Material (SRM)

The SRM analyzed in connection with mercury analysis of invertebrate samples was DORM-2 certified by the National Research Council of Canada.  DORM-2 is a dogfish muscle reference material.  The SRM analyzed in connection ICP and ICPMS metals analysis was DOLT-1, also certified by the National Research Council of Canada.  DOLT-1 is a dogfish liver reference material.  SRM recoveries outside the recommended 80 to 120% QC limits were reported for chromium (122%) and zinc (76%).

Matrix Spikes and Matrix Spike Duplicates

MS/MSD recoveries outside the recommended QC limits of 80 to 120% and/or exceeding the 20% RPD QC limit are summarized in the following table.

| Metal | MS Recovery | MSD Recovery | RPD |
| --- | --- | --- | --- |
| Antimony | 63% | 66% | 5 |
| Arsenic | 118% | 138% | 16 |
| Chromium | 5% | 292% | 193 |
| Copper | -18% | 466% | 216 |
| Nickel | 91% | 363% | 120 |
| Zinc | 51% | 91% | 56 |

Substantial matrix interferences were observed during ICP analysis of the invertebrate samples.  The matrix interferences necessitated the analysis of the samples in a diluted state.  Most affected by the interference was sample 11444-1 on which the MS and MSD were performed.  It also appears that the MSD for ICPMS analysis failed as evidenced by the extremely high recoveries for arsenic, copper and nickel.

Laboratory Duplicate Samples

The reported RPDs for all ICP and ICPMS laboratory duplicate results were less than the QC limit of 20%.  A mercury laboratory duplicate was not analyzed specifically on the invertebrate matrix, however, three laboratory duplicates were analyzed in the same QC batch which included clams and English sole.  The reported RPDs for the three mercury laboratory duplicate results were all less than the QC limit of 20%.

ORGANICS

Completeness

Organics data are reported for samples 11193-1, 11193-2, 11444-1, and 11444-2. These samples were analyzed for base/neutral/acid extractable semivolatile compounds (BNAs), polychlorinated biphenyls (PCBs), tributyltin, and percent lipids in association with the complete set of QC samples outlined in Table 2.

Methods

BNA analysis was performed in accordance with EPA method 8270 (SW-846).  PCB analysis was performed in accordance with EPA method 8080 (SW-846). Tributyltin analysis was performed in accordance with KCEL methodology which follows guidance specified in several papers (Unger, et.al., 1986; Krone, et. al., NOAA, 1989; and the Washington State Department of Ecology Manchester Laboratory Standard Operating Procedure).  Tributyltin analytical results were quantified using GC/MS in the selected ion monitoring (SIM) mode.  Percent lipid determination is performed using methodology found in the KCEL standard operating procedure OR 07-01-001.  This methodology involves a gravimetric determination of the percent lipid in the sample subsequent to extraction using methylene chloride and acetone.

Target List

The reported BNA target list includes all compounds on the EPA priority pollutant list in addition to caffeine, coprostanol, and several other BNA compounds.  Reported PCB data include Aroclors 1016, 1221, 1232, 1242, 1248, 1254, and 1260.

Detection Limits, Units, and Significant Figures

Data are reported in accordance with laboratory policy at the time the data were generated.  A positive result and/or MDL and RDL have been reported for all organic compounds with the exception of percent lipids.  Sample results for BNAs, PCBs and tributyltin are reported in units of mg/Kg on a wet weight basis.  Lipid data are reported in percent.  Data are reported to three significant figures for results greater than the RDL and two significant figures for results equal to or less than the RDL.  For results reported with less than two or three significant figures, significant zeroes are implied.

Storage Conditions and Holding Times

Sample storage conditions and holding times have been evaluated using guidelines established during the recent update of the Puget Sound Protocols (Recommended Guidelines for Sampling Marine Sediment, Water Column, and Tissue in Puget Sound - PSEP, 1996).  The criteria used to evaluate storage conditions and holding times for organics analyses are listed in the table below.

| Parameter | Frozen Holding Time | Refrigerated Holding Time |
| --- | --- | --- |
| BNAs | 1 Year to Extract 40 Days to Analyze | 14 Days to Extract 40 Days to Analyze |
| PCBs | 1 Year to Extract 40 Days to Analyze | 14 Days to Extract 40 Days to Analyze |
| Percent Lipids | 1 Year to Extract 40 Days to Analyze | 14 Days to Extract 40 Days to Analyze |
| Tributyltin | 1 Year to Extract 40 Days to Analyze | 14 Days to Extract 40 Days to Analyze |

All tissue samples in this data submission were stored frozen, both prior and subsequent to processing and homogenization.  All analyses were completed within holding times.

Base/Neutral/Acid (BNA) Extractable Semivolatile Compounds

Method Blanks

All BNA method blank results were less than the MDL.

Spike Blanks

BNA spike blank recoveries outside the recommended 50 to 150% QC limits are summarized in the following table.

| Compound | Recovery |
| --- | --- |
| 2,4-Dimethylphenol | 20% |
| 2,4-Dinitrophenol | 33% |
| Pentachlorophenol | 43% |
| Benzidine | 0% |
| Benzo(g,h,i)perylene | 47% |
| Benzoic Acid | 0% |
| 4-Chloroaniline | 38% |

Surrogate Recoveries

Eight surrogate compounds are added to samples during BNA analysis.  The four base/neutral surrogate compounds are 2-fluorobiphenyl, d14-terphenyl, d4-1,2-dichlorobenzene, and d5-nitrobenzene.  The four acid surrogate compounds are 2,4,6-tribromophenol, 2-fluorophenol, d4-2-chlorophenol, and d5-phenol.  BNA spike blank recoveries outside the recommended 50 to 150% QC limits are summarized in the following table.

| Surrogate | Sample | Recovery |
| --- | --- | --- |
| 2-Fluorophenol | 11193-1 | 31% |
| 2-Fluorophenol | 11444-2 | 49% |
| d14-Terphenyl | 11193-1 | 197% |
| d14-Terphenyl | 11193-2 | 201% |
| d4-1,2-Dichlorobenzene | 11193-1 | 40% |
| d5-Nitrobenzene | 11444-1 | 0% |
| d5-Nitrobenzene | 11444-2 | 0% |
| d5-Phenol | 11193-1 | 47% |

BNA target compounds that may be affected by the low recovery of surrogates d4-1,2-dichlorobenzene, d5-phenol and 2-fluorophenol include phenol, bis(2-chloroethyl)ether, 2-chlorophenol, 2-methylphenol, 1,2-dichlorobenzene, 1,3-dichlorobenzene, 1,4-dichlorobenzene, 2,2’-oxybis-(1-chloropropane), 4-methylphenol, N-nitroso-di-n-propylamine, and hexachloroethane.  Associated sample data for these compounds may be biased low.

BNA  target compounds that may be affected by the low recovery of surrogate d5-nitrobenzene include nitrobenzene, isophorone, 2-nitrophenol, 2,4-dimethylphenol, bis(2-chloroethoxy)methane, 2,4-dichlorophenol, 1,2,4-trichlorobenzene, naphthalene, 4-chloroaniline, hexachlorobutadiene, 4-chloro-3-methylphenol, and 2-methylnaphthalene. Associated sample data for these compounds may be biased low.

BNA  target compounds that may be affected by the high recovery of surrogate d14-terphenyl include pyrene, butyl benzyl phthalate, 3,3’-dichlorobendizine, benzo(a)anthracene, bis(2-ethylhexyl) phthalate, and chrysene.  Associated sample data for these compounds may be biased high.

Matrix Spikes (MS) and Matrix Spike Duplicates (MSD)

BNA MS/MSD recoveries outside the recommended 50 to 150% QC limits are summarized in the following table.

| Compound | MS Recovery | MSD Recovery |
| --- | --- | --- |
| Bis(2-chloroethyl)ether | 59% | 34% |
| 2-Chlorophenol | 65% | 49% |
| 1,3-Dichlorobenzene | 59% | 30% |
| 1,4-Dichlorobenzene | 77% | 40% |
| 1,2-Dichlorobenzene | 73% | 39% |
| Bis(2-chloroisopropyl)ether | 33% | 17% |
| N-Nitroso-di-N-propylamine | 0% | 0% |
| Hexachloroethane | 192% | 138% |
| 4-Chloro-3-methylphenol | 0% | 0% |
| Hexachlorocyclopentadiene | 0% | 0% |
| 2,4-Dinitrophenol | 32% | 0% |
| 4-Nitrophenol | 0% | 0% |
| 4,6-Dinitro-o-cresol | 28% | 13% |
| N-Nitrosodiphenylamine | 38% | 32% |
| 1,2-Diphenylhydrazine | 32% | 24% |
| 4-Bromophenyl phenyl ether | 51% | 42% |
| Hexachlorobenzene | 54% | 40% |
| Pentachlorophenol | 16% | 24% |
| Di-N-butyl phthalate | 43% | 34% |
| Benzidine | 0% | 0% |
| Pyrene | 202% | 140% |
| Benzyl butyl phthalate | 0% | 0% |
| 3,3’-Dichlorobenzidine | 36% | 27% |
| Benzo(g,h,i)perylene | 69% | 23% |
| Aniline | 0% | 0% |
| Benzoic Acid | 53% | 44% |
| 4-Chloroaniline | 33% | 37% |
| Carbazole | 52% | 39% |
| Coprostanol | 46% | 36% |

The RPDs for MS/MSD results were all less than the QC limit of 100%.  It should be noted that recoveries of five of the eight surrogates in the MSD sample were less than the lower QC limit of 50%.

Standard Reference Material (SRM)

An appropriate SRM was not available to analyze with the invertebrate tissue samples.

Laboratory Duplicate Samples

The RPDs for BNA laboratory duplicate sample results were all less than the 100% QC limit.

Polychlorinated Biphenyl (PCB) Compounds

Method Blanks

All PCB method blank results were less than the MDL.

Spike Blanks

The PCB spike blank recovery (Aroclor 1260) was within the recommended 50 to 150% QC limits.

Surrogate Recoveries

Two surrogate compounds are added to samples during PCB analysis, decachlorobiphenyl (DCB) and 2,4,5,6-tetrachloro-m-xylene (TCX).  Sample analytical data may be used unqualified if at least one of the PCB surrogate compounds is within the 50 to 150% QC limits.  At least one of the PCB surrogate compounds was within QC limits for all samples in this data submission with the exception of 11444-2.  The surrogate recoveries for sample 11444-2 were 17% for DCB and 11% for TCX.

Matrix Spikes (MS) and Matrix Spike Duplicates (MSD)

Aroclor 1260 is used as the spiking compound during MS/MSD analyses.  The MS/MSD recoveries were within the recommended 50 to 150% QC limits.

Laboratory Duplicate Samples

The RPDs for the PCB laboratory duplicate sample results were all less than the 100% QC limit.

Percent Lipid Analysis

Method Blanks

Lipids were detected in the method blank at a concentration of 0.03%.  All sample percent lipid concentrations were greater than 10 times the blank concentration

Laboratory Duplicate Samples

The RPDs for percent lipids laboratory duplicate sample results were all less than the 100% QC limit.

Tributyltin

Method Blanks

The tributyltin method blank result was less than the MDL.

Spike Blanks

All tributyltin spike blank recoveries were within the recommended 50 to 150% QC limits.

Internal Standards

The internal standards (similar to surrogates) used for this analysis are tripropyltin and tripentyltin.  Recoveries for both internal standards were within QC limits for all samples in this data submission.

Matrix Spikes (MS) and Matrix Spike Duplicates (MSD)

The MS/MSD recoveries were within the recommended 50 to 150% QC limits.

Laboratory Duplicate Samples

The RPD for tributyltin laboratory duplicate sample results was less than the 100% QC limit.

TABLE 1

DUWAMISH ESTUARY/ELLIOTT BAY INVERTEBRATE STUDY

TISSUE CHEMISTRY SAMPLE INVENTORY

| Sample | Locator | Mercury | Metals (ICP) | Metals (ICPMS) | PCBs | BNAs | Tributyltin | Percent Lipids |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 11193-1 | Duwamish River | X | X | X | X | X | X | X |
| 11193-2 | Duwamish River | X | X | X | X | X | X | X |
| 11444-1 | Nisqually Delta | X | X | X | X | X | X | X |
| 11444-2 | Nisqually Delta | X | X | X | X | X | X | X |

TABLE 2

QC SAMPLE FREQUENCY FOR TISSUE CHEMISTRY PARAMETERS

| Parameter | Method Blank | Spike Blank | Matrix Spike | Matrix Spike Dup. | Laboratory Dup. | SRM | Surrogate |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Mercury | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | No |
| Metals | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | No |
| BNAs | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | Yes |
| PCBs | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | No | Yes |
| Tributyltin | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | 1 Per Batch | No | Yes |
| Percent Lipids | 1 Per Batch | No | No | No | 1 Per Batch | No | No |

METALS QC DATA

ORGANICS QC DATA